The DG5MK IV Meter – An Accurate Antenna Analyzer

QEXfiles Addendum

Michael Knitter, DG5MK

This material supplements the

“Putting It All Together – The DG5MK IV-Meter Implementation”

section of the main article.

The key to the QEXfiles content is:

IV_Meter_V1.1_Bridge_Head_15.07.2016_Production.brd — board layout v1.1

IV_Meter_V1.1_Bridge_Head_15.07.2016_Production.sch — board layout v1.1

P_16_circuit_diagram_V1.1.png —schematic diagram of the IV meter

P_18_board_layout_v1.1.png — board layout v1.1

P_21_early_prototype_if_band.jpg — early prototype if bands

P_22_assembly_layout_v1.1.png — assembly layout v1.1

P_22_Control_Interfaces_V02.png — Control Interfaces V02

P_24.pdf — Selected sample measurement results.

The Software folder contains the code file IV_Meter_HEX_020b_16X2_07_2016.hex

Referring to the schematic diagram in P_16_circuit_diagram_V1.1.png, and starting on the left side in the middle and lower part, there is the microcontroller and LCD displays (a graphical display is shown for future use). Both display alternatives are shown since the PCB is designed to work with both versions.

The standard two-line display is connected to the microcontroller via 7 I/O-ports, and the (future) graphical display will need an SPI bus interface with some additional connections for reset, and so on. Since SPI is quiet common for graphical displays, the bus and additional connections will be available at an external connector as well. This allows to use other graphical displays, after modifying the software.

The MCP3911 ADC is in the middle of P_16_circuit_diagram_V1.1.png. SPI bus is connected to the microcontroller, along with an interrupt line. The ADC signals when a new sample is available, and alerts the microcontroller to fill the sample buffer.

At the top right, the anti-aliasing low pass is between the ADC and the 74LVC4066 mixer. The voltage dividers on above the mixer are used to generate an operating point of half the supply voltage. The bridge test head on the very top right side is ac coupled and includes the additional resistors to provide the 50  input impedance to the DUT.

On the right side in the middle, the SI5351 serves as the signal and clock generator for the overall circuit. It is connected to the microcontroller via an I2C bus. This bus connects the EEPROM to the microcontroller, at the right bottom of the schematic. The I2C is also available on an external connector for future use.

The rest of the circuit diagram deals with power supply generation, decoupling, bypassing, and so on. V1.0 of the IV-Meter had LDO regulators for nearly every functional block, as a conservative measure. Further tests showed that they can all be omitted with minimal impact to the final accuracy of the IV-Meter. This not only saves cost, but also decreases power consumption, since all components can be driven from one voltage source without additional loss by LDOs.

At the top left, an MCP1640 is used to generate the needed main voltage (3.3 V) from the batteries. A second MCP1640 is needed for the two-line display since an additional 5  V supply is required.

There are also a many 10  resistors on the bus lines to prevent spikes and unwanted oscillations. It was not tested if these can be omitted.

The bill of material is below. Some component numbers are intentionally missing; they are reserved for the future graphical display version.

The Bill of Materials is below.

| Parts | Qty GLCD | Qty 16x2 | Value | Package | Description |
| --- | --- | --- | --- | --- | --- |
| R16, R19, R44 | 3 | 3 | 0 | R0805 | RESISTOR, European symbol |
| R20, R21, R22, R24, R25, R26, R36, R37, R38, R40 | 10 | 10 | 10 | R0805 | RESISTOR, European symbol |
| R23 | 1 | 1 | 24 | R0805 | RESISTOR, European symbol |
| R41, R42, R43 | 3 | 3 | 49.9 | R0805 | RESISTOR, European symbol |
| R8 | 0 | 1 | 68 | R1206 | RESISTOR, European symbol |
| R27 | 1 | 1 | 100 | R0805 | RESISTOR, European symbol |
| R45, R46, R47, R48 | 4 | 4 | 240 | R0805 | RESISTOR, European symbol |
| R13, R14, R32, R33, R34, R35 | 6 | 6 | 1k | R0805 | RESISTOR, European symbol |
| R1, R9, R10, R15, R28, R29, R30, R31 | 7 | 8 | 10k | R0805 | RESISTOR, European symbol |
| R12 | 1 | 1 | 309k | R0805 | RESISTOR, European symbol |
| R6 | 1 | 1 | 330k | R0805 | RESISTOR, European symbol |
| R7 | 1 | 1 | 560k | R0805 | RESISTOR, European symbol |
| R11 | 0 | 1 | 976k | R0805 | RESISTOR, European symbol |
| R17, R18, R39 | 0 | 0 | 9999k | R0805 | RESISTOR, European symbol |
| R2 | 0 | 1 | 5k | RTRIM64P | Trimm resistor |
| C16, C17 | 2 | 2 | 5.6p | C0805 | CAPACITOR, European symbol |
| C14, C15, C32, C33 | 4 | 4 | 10n | C0805 | CAPACITOR, European symbol |
| C2, C13, C18, C21, C23, C25, C26, C28, C29, C34, C35, C36, C37, C38 | 13 | 14 | 100n | C0805 | CAPACITOR, European symbol |
| C30, C31 | 11 | 2 | 1u | C0805 | CAPACITOR, European symbol |
| C1, C3, C19, C20, C27, C39, C40 | 5 | 7 | 22u | C0805 | CAPACITOR, European symbol |
| L1, L4 | 1 | 2 | 4.7uH | L1812 | INDUCTOR, Fastron 1616FPS-4R7M-01 |
| L3 | 1 | 1 | 600 | C1206 | INDUCTOR, Ferrit Murata BLM31AJ 601 SN1L |
| Q2 | 1 | 1 | 27MHz | SM49 | CRYSTAL HC49 SMD |
| IC1, IC3 | 1 | 2 | MCP1640 | SOT95P270X145-6N | Synchronous Boost Regulator |
| IC10 | 1 | 1 | 24FC256 | SO-08 | Serial EEPROM, Microchip |
| IC11 | 1 | 1 | SI5351A | MSOP10 | I2C clock generator, SI |
| IC2 | 1 | 1 | DSPIC33FJ128GP802-I/SO | SOIC-28 | 16-bit Digital Signal Controller |
| IC8 | 1 | 1 | MCP3911A0-E/SS | SOP20 | Two-Channel Analog Front End,SSOP-20 |
| IC9 | 1 | 1 | 74LVC4066 | SOIC-14 | Quad bilateral switches |
| Q1 | 1 | 1 | BC848BSMD | SOT23 | NPN Transistor |
| DIS1 | 0 | 1 | 16X2_R2 | TUXGR_16X2_R2 | LCD display 16x2 characters |
| S1, S2, S3 | 3 | 3 | DT6 | D6R | ITT SWITCH |
| JP1 | 1 | 1 | U_Bat_3V | 1X02 | PIN HEADER |
| JP3, JP4 | 2 | 2 | PGM, SPI LCD | 1X06 | PIN HEADER |
| JP6, JP7 | 2 | 2 | BUTTONS, I2C | 1X04 | PIN HEADER |
| X1 | 1 | 1 | SMA | J502-ND-142-0701-801/806 | SMA 50 Ohm End Launch Jack |
| PCB 1 | 1 | 1 | PCB |  | PCB V1.1 |
| Totals | 93 | 94 |  |  |  |

Image P_18_board_layout_v1.1.png shows the Version 1.1 of the PCB layout, before generating the ground planes. The buttons and the display are assembled from the backside of the board. All other components go to the top side. Most SMD components are 0805 size and can easily be assembled by hand. The most challenging parts are the MCP3911 and the SI5351 with their small outlines. Internet searches provide some good tutorials on how to best solder such components.

============================================